## Annex 1
List of all items required |  |  | To be filled by the bidder
Item # | Description | Qty. required | Manufacturer | Model | Country of Origin | FDA/CE | Warranty (Years) | Unit Price (USD) | Total Price (USD) | Delivery Time | Remarks
1 | Portable Ultrasound | 1

## Item 1
Item # 1 | Compliance (Yes/No) | Deviations
Portable Ultrasound Scanner (General Purpose) (Qty.1)
Manufacturer (Please specify)
Model (Please specify)
FDA and /or European CE Approval (Please attach it with the offer)
Country of Origin:
Delivery Time (Please specify)
Unit price (USD)
Total price (USD)
Description of Function
Portable laptop-style ultrasound machines suitable for mobile clinics with a high patient volume.
Technical Specifications
Not Less than 15 " (B/W) & color LED Display High Resolution
Ready to use with Linear, Curved Array, Phased Array
B-mode, M-mode, B/M modes
Dual Image Feature
Zoom. Reverse, Invert image features
Depth selectable 3.5 up to 20 cm
Full Keyboard and track ball
Measurements and calculations to provide for Distance, Length, Area, Circumference.
Application packages for OB,GYN and Abdomenal.
Rechargeable Battery Bult - in
100-240 VAC - Frequency 50/60 Hz
Ready to work with software license and all needed accessories
Included Accessories
One Convex Probe 3 - 5 MHz
Hard Carrying Case is included
B/W Video thermal printer
Paper Rolls for printer (Qty. 5)
Gel pack (Qty.2)
Other Requirements
Warranty at least two years
On-site Application Training
Complete with all accessories and consumables needed to work completely as specified.
Service and Operating Manual Included
A valid Manufacturer Authorization Letter is included
References for the offered model (please mention installed base for the offered model)
To be priced separately
Color Flow, Pulse Wave, Power Doppler
Linear probe
Phased Array Probe
Original Trolley with storage area
Paper Roll for printer